CLASP (Continuous lifestyle awareness through sweat platform): A novel sensor for simultaneous detection of alcohol and glucose from passive perspired sweat
Ashlesha Bhidea, Sriram Muthukumarb, and Shalini Prasada*
aDepartment of Bioengineering, University of Texas at Dallas, Richardson, TX 75080
bEnLiSense LLC, 1813 Audubon Pond Way, Allen, TX 75013
*Corresponding author, Email: shalini.prasad@utdallas.edu

Supplementary information

200nm

Figure S1: Scanning electron microscope (SEM) image of ZnO deposited on the surface of polyamide membrane.

Figure S2: Energy dispersive X-ray spectroscopic (EDAX) spectrum of ZnO deposited on the surface of polyamide membrane.

Figure S3: Chronoamperograms for combinatorial sensing of alcohol and glucose in synthetic sweat pH 6 on an alcohol sensor.

Figure S4: Chronoamperograms for combinatorial sensing of alcohol and glucose in synthetic sweat pH 6 on a glucose sensor.

Figure S5: Standard Randles circuit representing the biosensor system.
Rs: Resistance of the solution/buffer.
RZnO: Resistance of the zinc oxide thin film.
CZnO: Capacitance of the zinc oxide thin film.
Cdl: Capacitance of the electrical double layer.
Rct: Charge transfer resistance.
Zw: Warburg impedance.
Table S1: Comparison of the proposed sensor with other alcohol and glucose sensors for continuous monitoring in bodily fluids
Substrate	Analyte and Buffer	Transduction / mechanism	Dynamic range	Advantages	References
Alcohol oxidase modified screen printed Prussian blue electrodes	Ethanol in human sweat	Amperometry	0-36mM	Non-invasive, wearable, real- time wireless on- body alcohol monitoring, uses iontophoresis to induce sweat	(Kim et al., 2016)
Bienzyme composite graphite-Teflon electrode	Ethanol in human sweat	Amperometry	5E-03- 0.6 gL-1	Determines the ethanol content equivalent to BAC content of the user at given time or over 8hrs in real-time, correlation of BAC and TAC.	(Gamella et al., 2014)
Pt wire transducer coated with PPD/AlOx-Chit/Nafion multilayer.	Ethanol in interstitial fluid (ISF)	Amperometry	0-80mM	Microneedle based sensor for continuous alcohol monitoring in ISF, efficacy of the sensor tested on an ex-vivo mice skin model, stable current response over 100min duration	(Mohan et al., 2017)
Gold doped graphene electrode	Glucose in human sweat	Electrochemical impedance spectroscopy	10M- 0.7mM	Combined detection of glucose, pH, and humidity, microneedles for drug delivery for hyperglycemia conditions	(Lee et al., 2016)
Enzyme based- Prussian blue electrodes	Glucose in human sweat	Chronoamperometry	0- 200M	Combined detection of glucose, lactate, and electrolytes in human sweat, on-body real-time monitoring of physiological conditions	(Gao et al., 2016)
Glucose oxidase modified Prussian blue electrodes	Glucose in human sweat	Amperometry	0- 2mM	Wearable sensor integrated into eyeglasses capable of real-time monitoring of sweat electrolytes and metabolites during exercise	(Sempionatto et al., 2017)
Enzyme immobilized gold- zinc oxide electrodes	Alcohol and glucose in human sweat	Electrochemical impedance spectroscopy and chronoamperometry	0.01- 100mg/dl (Alcohol) / 0.01- 50mg/dl (Glucose)	Non-invasive, dynamic combinatorial detection of alcohol and glucose in human sweat, detection possible using ultra- low volumes (1-5µL) of passive sweat	This work

Table S2: Comparison of sensor metrics computed for biomolecules spiked in synthetic sweat pH 4 using EIS and CA techniques
Technique	EIS		Chronoamperometry
Molecule	Alcohol	Glucose	Alcohol	Glucose
LOD	0.01 mg/dl	0.01 mg/dl	0.01 mg/dl	0.01 mg/dl
Dynamic range	0.01- 100mg/dl	0.01- 50mg/dl	0.01- 100mg/dl	0.01- 50mg/dl
Sensitivity	10.5 Ω/mg dl-1	19.2 Ω/mg dl-1	0.056µA/mg dl-1	0.12µA/mg dl-1

Table S3: Comparison of sensor metrics computed for biomolecules spiked in synthetic sweat pH 6 using EIS and CA techniques
Technique	EIS		Chronoamperometry
Molecule	Alcohol	Glucose	Alcohol	Glucose
LOD	0.01 mg/dl	0.01 mg/dl	0.01 mg/dl	0.01 mg/dl
Dynamic range	0.01- 100mg/dl	0.01- 50mg/dl	0.01- 100mg/dl	0.01- 50mg/dl
Sensitivity	12.8 Ω/mg dl-1	38.8 Ω/mg dl-1	0.037µA/mg dl-1	0.07µA/mg dl-1

Table S4: Comparison of sensor metrics computed for biomolecules spiked in synthetic sweat pH 8 using EIS and CA techniques
Technique	EIS		Chronoamperometry
Molecule	Alcohol	Glucose	Alcohol	Glucose
LOD	0.01 mg/dl	0.01 mg/dl	0.01 mg/dl	0.01 mg/dl
Dynamic range	0.01- 100mg/dl	0.01- 50mg/dl	0.01- 100mg/dl	0.01- 50mg/dl
Sensitivity	14.3 Ω/mg dl-1	27.2 Ω/mg dl-1	0.032µA/mg dl-1	0.076µA/mg dl-1

Table S5: Comparison of sensor metrics computed for biomolecules spiked in human sweat using EIS and CA techniques

Technique	EIS		Chronoamperometry
Molecule	Alcohol	Glucose	Alcohol	Glucose
LOD	0.1 mg/dl	0.1 mg/dl	0.01 mg/dl	0.01 mg/dl
Dynamic range	0.1- 100mg/dl	0.1- 50mg/dl	0.01- 100mg/dl	0.01- 50mg/dl
Sensitivity	15.4 Ω/mg dl-1	14.2 Ω/mg dl-1	0.02µA/mg dl-1	0.048µA/mg dl-1

Computation of specific signal threshold (SST)
1. Percentage change in impedance =
2. SST for EIS based CDR’s - The specific signal threshold (SST) is three times the ratio between the standard deviation of impedance response measured at the baseline step the average impedance response measured at the baseline step. Percentage SST is computed using the expression (300*σbaseline/µbaseline).
3. SST for CA based CDR’s - The SST is calculated for an SNR of 3 using the expression (µ∆z - 3*σ∆z) where ∆z is the difference in the current measured between the enzyme immobilization step and the baseline step.

References
1. Gamella, M., Campuzano, S., Manso, J., Rivera, G.G., Lopez-Colino, F., Reviejo, A.J., Pingarron, J.M., 2014. Anal. Chim. Acta. 806, 1-7.
2. Gao, W., Emaminejad, S., Nyein, H.Y.Y., Challa, S., Chen, K., Peck, A., Fahad, H.M., Ota, H., Shiraki, H., Kiriya, D., Lien, D., Brooks, G.A., Davis, R.W., Javey, A., 2016. Nature. 529, 509-514.
3. Kim, J., Jeerapan, I., Imani, S., Cho, T.N., Bandodkar, A., Cinti, S., Mercier, P.P., Wang, J., 2016. ACS Sens. 1, 1011-1019.
4. Lee, H., Choi, T.K., Lee, Y.B., Cho, H.R., Ghaffari, R., Wang, L., Choi, H.J., Chung, T.D., Lu, N., Hyeon, T., Choi, S.H., Kim, D., 2016. Nature Nanotechnology. 11, 566.
5. Mohan, A.M.V., Windmiller, J.R., Mishra, R.K., Wang, J., 2017. Biosensors and Bioelectronics. 91, 574-579.
6. Sempionatto, J.R., Nakagawa, T., Pavinatto, A., Mensah, S.T., Imani, S., Mercier, P., Wang, J., 2017. Lab Chip. 17, 1834-1842.